## PSA plant + container + booster
Vendor
Model
Product | Category | Product specifications (adapted from WHO 2020) | Meet requirement? (Y/N) | Type of supporting documents (select from dropdown) | Page no. in supporting documents | Comments
PSA plant + container + filling station/booster compressor | Overview | Pressure Swing Adsorption (PSA) technology
• Rated to up to 40 Nm3/hr
• Molecular sieve beds comprising zeolites
• Rated for medical application
Cylinder filling station/ramp, booster compressor to align with plant’s full output (up to 40 Nm3/hr) to use all capacity.
Containerized solution
Duplex configuration of any/all components favored.
| Oxygen output (purity & pressure) | 93±3% with continuous output, produced from ambient air
Inbuilt analyser: +/- 1% accuracy
Output pressure: between 300-600 kPa (3-6 bar, 44 - 87 psi).
| Air intake Compressor & dryer: | • Either oil-free or filtered oil-injected or oil-lubricated rotary screw-type
• Maximum output pressure of 750 kPa/7.5 bar
• External refrigerant air dryer
• Soft-start variable speed drive (VSD) compressor (optional)
| Filters | • Air intake side: Filtration assembly, comprising:
o Particulate filter (>5 micron)
o Coalescing filter (0.01 micron)
o Carbon absorption filter (Coal tower, if applicable)
• Product/outlet side: Bacterial outlet (product) filter
| Power requirments | Compressor operations: 380 volts /50 Hz (3 Phase)
Control system operations: 220 VAC / 50 Hz (Single phase)
• Electrical requirements must be locally compatible, plug Type E
• Electrical protection by resettable circuit breakers or replaceable fuses, fitted in both neutral and live lines.
| System operations | Program logic-controlled System
Touch-screen control panel in French
• Integrated O2 purity analyzer and monitor for medical application, paramagnetic type (or indicate otherwise)
• Displays:
o Purity Oxygen concentration [%], with +/- 1% accuracy),
o production trending (Nm3/hr),
o output pressure,
o overall system status & time to next service
o cumulative operating hours
• All displayed in SI units
• Audible and visual alarms for:
o elevated temperature
o low purity
o low (<3bar) or high (>6bar) pressure
o power and/or system failure
o air dryer PDP (>3°C)
o secondary or reserve source (if applicable)
• Remote monitoring feature (optional)
• Automatic drain on air receiver tank
• Function of purge of low concentration of oxygen
| Cylinder filling station | Oxygen boosting compressor (oil-free):
• 15,000 kPa output pressure (150 bar)
Cylinder filling ramp with 8 (or more) connection points
• bullnose valve connection NEN "Ri2" (Oxygen)
Cylinders to have secure fastening (e.g., chain, bracket)
Purge compressor/vacuum to empty cylinders
| Mobility / installation ready | Containerized solution, pre-configured with requisite intake, ventilation, HVAC and exhaust vents and louvres.
| Spares & consumables | • Three (3) year spare parts kits as per recommended preventative maintenance programme clearly defined in a disaggregated list comprising part numbers, descriptions, and unit cost as well as indicating brand/model specifics (e.g. for circuit breaker, printed circuit board, sieve beds, compressor components, valves, wheels, motor capacitor, analyzer, etc.) by the manufacturer
• Set of inlet filters and outlet bacteria filters for three (3) year operation.
| Environmental considerations | Capable of supplying the specified oxygen concentration continuously:
• In ambient temperature from 10–40 °C,
• At relative humidity from 15-95%, preferably simultaneously.
• At elevation: 1,500 m above sea level
| Installation, testing and commissioning | Pre-shipment: Certificate of quality, calibration, and inspection
On-site:
• Installation to include training on start-up, use, and maintenance
• Inspection, testing and commission should be done before handover.
Verification to include:
• Oxygen purity, pressure levels
• Electrical system functionality, alarms functionality
• Automatic switchover to secondary/reserve (if applicable)
If possible, conformity of installation shall be verified by a certified third party. Otherwise, alternative verification process to be agreed upon between vendor and owner.
| Documentation | Hard and soft copies, in English AND French of the following:
• Certificate of quality, calibration, and inspection
• User manual, detailing
o Protocols for operations
o List of procedures and equipment required for cleaning
o List of procedures and equipment for preventive maintenance
o Troubleshooting chart
• Contact details of manufacturer, and authorized distributors (if applicable), and local service agents
• Product and model reference (brochure)
| Warranty | All components of PSA oxygen plant system should have at least 3 years’ warranty period after commissioning
The supplier must ensure the availability of spare parts for at least 8 years
| Regulatory and standards | Regulatory approval: Class C (GHTF Rule 11); FDA Class II (USA); Class IIb (EU) under MDR; Class IIa (Australia); Class II (Canada).
Standards: the following certificates given by a certified third-party for the system proposed (or equivalent thereof):
General:
• Certified Quality Management Systems (ISO 13485, ISO 9001)
Component-specific:
• ISO 7396-1: Medical gas pipeline systems — Part 1: Pipeline systems for compressed medical gases and vacuum.
• ISO 8573-1: Compressed air – Part 1: Contaminants and purity classes.
• ISO 8573-2: Compressed air – contaminant measurement – Part 2: Oil aerosol content.
• ISO 8573-4: Compressed air – contaminant measurement – Part 4: particle content.
• ISO 5011: Inlet air cleaning equipment for internal combustion engines and compressors – performance testing.
• ISO 21969: High pressure flexible connections for use with medical gas systems.
All pressurized vessels to be:
o designed according to PED or ASME VIII, or equivalent;
o certified PED or ASME III, or equivalent;
Cleaned according to ISO 15001, ASTM G93, or equivalent.
| Service Agreement | • 36 months • Provide o training schedule o training content o schedules for planned preventive maintenance (PPM) o define corrective maintenance, including repair agreements, delineating what falls out of warranty o provide time for response o provide time for availability of critical spares  origin and/or warehouse location of spares
Provide:
o training schedule
o training content
o schedules for planned preventive maintenance (PPM)
o define corrective maintenance, including repair agreements, delineating what falls out of warranty
o provide time for response
o provide time for availability of critical spares
• origin and/or warehouse location of spares

## Cylinders
Vendor
Model
Product | Category | Product specifications (adapted from WHO 2020) | Meet requirement? (Y/N) | Type of supporting documents (select from dropdown) | Page no. in supporting documents | Comments
Cylinders | DESCRIPTION | A container designed as a refillable cylinder used to hold compressed medical oxygen (O2) under safe conditions at high pressure (e.g. 50–200 bar). It is typically filled with oxygen when delivered from the gas supplier and includes a valve stem, an opening/closing valve, and will be graded according to size (capacity) and colour-coded to denote oxygen content. The cylinder may be made of steel, aluminium (Al) or other ferrous or non-ferrous materials and must be used together with a pressure regulator in order to release the oxygen at the correct working pressure. (SOURCE: GMDN 47225)
| Technical | Refillable cylinders for compressed oxygen gas for medical use.
Material: CGA approved seamless steel/aluminum alloy/carbon fibre/other composite body
Size "B50" (50L water capacity)
Nominal / working pressure: 137-150 bar
Primary valve: Bullnose NEN "Ri2" (Oxygen)
Valve assembly material: brass
Valve handle - hand-wheel
Colour coding: according to ISO 32 (black body, white shoulders)
Supplied with valve cap
Stamped onto shoulder:
· Nominal and test pressures
· Cylinder capacity (water litres)
· Weight (kg)
· Date of manufacture
· Serial number
· Company and country of origin
· transport rating (DOT3AA or UN stamp)
· inspection agency stamp aligning with regulatory requirement
Environmental:
· Suitable for continuous operation in ambient temperature of at least 5–45 °C, relative humidity of at least 15–90% non condensing
· Capable of being stored in ambient temperature of at least 5–50 °C, relative humidity of at least 15–95% non condensing
| Warranty | 5 years recommended
| QMS | ISO 9001 (General QMS)
| Regulatory | CE indicated by pi, "p" (Cylinders fall under TPED) or US DOT
· applies to both cylinders AND valves
| Product performance standards Where testing indicated, must be done by an accredited laboratory - documentation of which to be included prior to shipment | ISO 10297: Gas cylinders – Cylinder valves – Specification and type testing
ISO 14246: Gas cylinders – Cylinder valves – Manufacturing tests and examinations
ISO 7866: Gas cylinders – Refillable seamless aluminium alloy gas cylinders – Design, construction and testing
ISO 9809-1: Gas cylinders – Refillable seamless steel gas cylinders – Design, construction and testing
ISO 13769: Gas cylinders — Stamp marking
ISO 13341: Gas cylinders – Fitting of valves to gas cylinders
ISO 15996: Gas cylinders – Residual pressure valves – Specification and type testing of cylinder valves incorporating residual pressure devices
ISO 11363-1: Gas cylinders — 17E and 25E taper threads for connection of valves to gas cylinders
ISO 15001: Anaesthetic and respiratory equipment – Compatibility with oxygen
| Regional/local standards | Country-specific and regional colour gas coding and other standards apply and must be listed.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol).
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.
Extra information required: number of units.

## Regulator, pressure (cylinder)
Vendor
Model
Product | Spec category | Specs, WHO-UNICEF derived | Meet requirement? (Y/N) | Type of supporting documents (select from dropdown) | Page no. in supporting documents | Comments
Pressure regulator | DESCRIPTION | A reduction valve designed to be attached through a gas-tight connector (e.g. a pin-index or bull-nosed screw thread connection) to the valve stem of an oxygen (O2) cylinder to lower high, variable gas pressure (e.g. 50–250 bar) to a lower, constant working pressure, typically 3–5 bar. It can be a single- or a dual-stage regulator, usually of a piston or diaphragm design. It will have a safety relief valve to avoid excessive pressure due to increased ambient temperatures, and it may have associated devices, usually a manometer or manometers, to display the available gas reserve of a gas cylinder and the working pressure. (SOURCE: GMDN 43438) High-pressure gas regulators are placed between the output port of the high-pressure gas container at one end and the patient breathing circuit at the other end.
| Technical | Pressure regulator inlet connection: compatible with cylinder bullnose valves NEN "Ri2" (Oxygen)
Outlet connection standard DISS, male
Steel/plated brass/aluminium casing, brass valve.
Nominal inlet pressure 13 700 kPa (137 bar, 1987 psi), maximum 20 000 kPa (200 bar, 2901 psi).
Outlet pressure 345 kPa (3.5 bar, 50 psi).
Integrated manometer, 0–20 000 kPa (0–200 bar, 0–2901 psi).
Safety over-pressure release valve.
Rated for use with medical oxygen (clearly indicated on manometer)
| Warranty | Min. 1 year (ideal 2)
| Standards | ISO 13485 (medical device QMS)
ISO 14971 (application of risk management)
| Regulatory | CE
FDA
| Product performance standards | ISO 10524 Pressure regulators for use with medical gases.
ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen.
ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## Flowmeter
Vendor
Model
Product | Spec category | Specs, WHO-UNICEF | Meet requirement? (Y/N) | Type of supporting documents (select from dropdown) | Page no. in supporting documents | Comments
Flowmeter, Thorpe tube, pressure compensated | Description | A device intended to measure and regulate the flow of a medical gas [e.g. oxygen (O2), carbon dioxide (CO2), nitrous oxide (N2O), helium/oxygen gas mixture (heliox), medical air] during various procedures (e.g. therapeutic administration, anaesthesia, insufflation during surgery). It consists of an upright tube containing a float, which rises and falls in relation to gas flow, and a distal valve (compensated flowmeter) to control gas flow rate. It will be calibrated to a specific medical gas and have a dedicated flow rate range, therefore some types may be dedicated to a specific patient group (e.g. neonate, infant, adult) or clinical use. (SOURCE: GMDN)
| Technical | Device suitable for use with medical oxygen
Thorpe tube flowmeter type, contains inlet and outlet port, a flow regulator, a valve and a clear measuring tube
Flowmeters to measure and regulate flow from an already pressure-reduced and regulated oxygen source to the patient or other medical device
Pressure compensated flowmeters, calibrated at 345–380 kPa (3.4–3.8 bar, 50–55 psi) inlet gauge pressure.
Max gauge inlet pressure 690 kPa (6.9 bar, 100 psi).
Flow adjustment knobs to have rough surface to prevent slipping.
Flowmaters calibrated to the following flow range, all metric (specify):
0-3.5 L/min, accuracy 10%, single taper graduations 0.25 L/min full range
0-15 L/min, accuracy 10%, dual taper graduations 0.5 L/min (0–5 L/min range) and 1 L/min (5 L/min – maximum range)
All minimum flowrates to be zero when fully closed
All graduations to be clearly visible for 270 degrees (most breadth for provider vantage points)
Inlet ports for cylinder use, connection to a cylinder pressure regulator:
Inlet connection of flowmeter to be DISS female
Outlets: DISS (9/16 inch-18) female to "Christmas tree" barbed 1/2 inch ID (M) tubing adaptor / hose connector, DISS humidifier adapter
Flowmeter material:
Column to be transparent, clear, shatter-reistant, medical-grade polymer (polypropylene, polycarbonate)
Hardware/valves: Brass/steel/aluminium
All materials in contact with oxygen certified for medical use.
Internal parts (e.g. valve, inlet filter if present), replaceable by user.
Environmental
Capable of being stored in ambient temperature of at least 5–50 °C, relative humidity of at least 15–95% non condensing.
Suitable for continuous operation in ambient temperature of at least 5–45 °C, relative humidity of at least 15–90% non condensing.
Specific requirements for altitude may be required depending on the installation site.
Disinfectable with hospital grade detergents.
| Warranty | Min. 1 year (ideal 2)
| QMS | ISO 13485 (medical device QMS)
ISO 14971 (application of risk management)
| Regulatory | CE
FDA
| Product performance standards | ISO 32 Gas cylinders for medical use — Marking for identification of content (or ANSI equivalent)
ISO 5359 Low-pressure hose assemblies for use with medical gases.
ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen.
ISO 15002 Flow-metering devices for connection to terminal units of medical gas pipeline systems.
ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.
ISO 18082 Anaesthetic and respiratory equipment – Dimensions of non-interchangeable screw-threaded (NIST) low-pressure connectors for medical gases.
ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## Humidifier, bubble (non-heated
Vendor
Model
Product | Spec category | Specs, WHO-UNICEF | Meet requirement? (Y/N) | Type of supporting documents (select from dropdown) | Page no. in supporting documents | Comments
Humidifier | Description | A device designed to prevent the drying of airway passages associated with the inhalation of oxygen (O2) by adding water vapour to the dry gas as it is passed through, or more seldom, over water. It typically consists of a graduated container (reservoir) for the water, a top piece that functions as a detachable lid (typically a screw lid with a gastight seal), and a tube that protrudes into the water to divert the gas below the water level. This device, commonly known as a bubble humidifier, does not heat the water. It has connectors: 1) one (e.g. a winged nut) that connects to an oxygen therapy flowmeter; and 2) one to which the patient tubing is connected. This is a reusable device. (SOURCE: GMDN 35113)
| Technical | Reusable humidifier for oxygen therapy and ventilation/anaesthesia inspiratory lines.
Non-heated humidifier - ambient temperature functionality.
Bubble-through humidification system.
Unbreakable or shatter resistant.
Transparent humidification bottle
Graduated, graduation shall show minimum and maximum water level.
Humidification chamber working volume at least 150 mL, not greater than 500 mL.
Detachable metal or rigid durable polymer cap with gas connectors.
Pressure relief safety valve, ≥ 14 kPa (0.1 bar, 2 psi) pressure rating.
DISS connectors for inlet.
6 mm barbed connector for outlet.
Flow rate capacity up to 15 L/min.
Must be capable of disinfection.
Materials, all to be certified for medical use:
·Cap and connectors made of brass/steel/other biocompatible metal or polymer
·Bottle and tubes made of polypropylene, polycarbonate or equivalent biocompatible plastic/polymer
·Pressure valve made of brass chromium plated or equivalent metal
Supplier must define decontamination procedure.
| Warranty | Min. 1 year (ideal 2)
| Standards | ISO 13485 (medical device QMS)
| Regulatory | CE
FDA
| Product performance standards | ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen.
ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.
ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications.
ISO 18562-1 Biocompatibility evaluation of breathing gas pathways in healthcare applications – Part 1: Evaluation and testing within a risk management process.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol).
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## RFI_MFLab (hidden)
| 关于产品 | [9 empty] | 关于贵司
产品类型 Product Type | 产品标准要求 Product Specifications | 产品型号（请列出贵公司符合左侧标准要求的产品，并邮件提供电子版产品用户手册） Product Model (s) | 产品排期（假设下周下订单，预计能够供的最早出货排期至何时（日期）？ Lead time if order placed next week (date) | 本次订单所需数量 Units Required | 公司当前可出售的数量 Units Available | 法规批准/认证Regulatory Approval/Certification |  | 产品性能标准Standards, for product performance |  | 产品单价（美金） Unit price (USD) | 制造商标准 Standards, for manufacturer |  | 预付款要求 Pre-payment requirement | 公司以前有向埃塞俄比亚出口的经验吗？Does company have previous export experience to Ethiopia? | 图片示例 Picture
[6 empty] | 标准 | 如果有，请填写 “X” | 标准 | 如果有，清填写 “X” |  | 标准 | 如果有，清填写 “X”
简单氧气气面罩 Simple Face Masks | 成人式： 医用级 | LB1110, Adult(LA), 5.5mm x 2m绿色管 |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X | 生产前付清 | 有
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
不重吸面罩 Oxygen Mask with reservoir bag | 成人式： 氧气浓度：支持80%至100％的氧气浓度 管长: 2m 使用标准接头；管路（氧气管线）不弯曲 材质：聚氯乙烯（PVC）； 透明；医用级 | LB1310, Adult(LA), 5.5mm x 2m绿色管 氧袋1000ml |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
文丘里面罩（氧浓度可调面罩） Venturi Mask | 成人式：能以8至10 L / min 的流速输送低至中等流量的氧气，特定氧浓度为30％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,31%(氧气流速6L/min±10%),白色，其他符合 |  | 300 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
| 成人式：能以10至12 L / min 的流速输送低至中等流量的氧气，特定氧浓度为40％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,40%(氧气流速8L/min±10%),粉色，其他符合 |  | 300 |  | CE | X | ISO 18190
[6 empty] | FDA | X | ISO18562-1 | X
[6 empty] | 出口许可证 | X | ISO/DIS 23368
| 成人式：能以12至15 L / min 的流速输送低至中等流量的氧气，特定氧浓度为60％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,60%(氧气流速8L/min±10%),红色，其他符合 |  | 300 |  | CE | X | ISO/DIS 17256 | X
[6 empty] | FDA | X | ISO 15223-1 | X
[6 empty] | 出口许可证 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 中等大小 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带, M |  | 130 |  | CE |  | ISO 18562-1:2017 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 17510:2015 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X |  |  |  | ISO14971 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 大码。 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带，L |  | 130
气管导管 Endotracheal Tube | 管： 尺码：7。 内径：7.0mm。 外径：9.3mm。 长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 袖口： 直径：24mm | LB5010C, 带囊，ID 7.0mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X |  | ISO14971 | X
| 管： 尺码：7.5 内径：7.5mm。长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 | LB5010C, 带囊，ID 7.5mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X