## 1a) ISOtainer (general)
Component | Specification | Meets requirement? (Yes/No/NA) | Type of supporting document (select from drop down) | File name as submitted | Page No. | Comments
Design Parameters | Horizontal construction, ISOtainer type
• Stainless steel or carbon steel outer shell
□ Affixed lifting lugs for use w/ clevis pins
• Stainless steel inner vessel (no aluminum)
• Stainless steel piping
• Vacuum pressure 0.05 bar (or better)
Maximum design pressure / test pressure (MAWP): 18 bar
Minimum / maximum operating temperatures:
• Inner vessel: -196°C /+20°C
• Outer shell: -30 °C /+50 °C
Sizing | 10,000L-20,000L LOX Capacity
Features | Operating:
• Fill coupling / type for LOX (spec type available)
□ Bottom and top-fill lines
• Pressure regulation
□ Pressure build-up/blow-down assembly
□ Pressure control for ISOtainer to microVIE filling (see tab 3)
• Overfill protection system, comprising
□ Fill line drain valve
□ Trycock vent valve
• Affixed ambient/passive vaporizer
Monitoring:
• Liquid level gauge (& type to be specified)
• Liquid level transmitter (& type to be specified)
Safety:
• Thermal relief to protect against temperature increase in ISOtainer
• Thermal relief valve to protect against temperature increase during ISOtainer to microVIE filling
• Reclosable safety valves to protect ISOtainer from overpressure: primary (set at operating pressure) and secondary (set at MAWP) (no bursting discs)
• Protect against overpressurization of microVIE to MAWP during ISOtainer to microVIE filling (see tab 3)
Additional offtake:
• Economizer circuit/condensing coil (preferred option)
• Discharge line for liquid offtake
□ Fitted with 5/8" x 45⁰ male flare for connection to transfer line for ISOtainer to microVIE filling (see tab 3)
Documentation | P&ID - generic
Sample databook, illustrative of what would be included (not limited to):
• Manufacturer's declaration of conformity
• 3rd party (e.g., Notified Body for TPED) certificate of conformity and test reports, certificates for safety valves
• Commissioning, operations, and maintenance manual
• Detailed stress / loading calculations
Export license
Training materials or services offered on start-up, use for ISOtainer to microVIE filling, and preventive/curative maintenance
Warranty | 15 years minimum
The supplier must ensure the availability of spare parts beyond the duration of the warranty
Regulatory | TPED 2010/35/EU or US DOT
QMS | ISO 9001: Quality Management Systems - Requirements
General design guidance (or equivalent) | BS EN 13530 - Cryogenic vessels. Large transportable vacuum insulated vessels.
Process or component specific standards (or equivalent) | EIGA 200/17: The Safe Design, Manufacture, Installation, Operation and Maintenance of Valves Used in Liquid Oxygen and Cold Gaseous Oxygen Systems
ISO 8501-1: Preparation of steel substrates before application of pains and related products – visual assessment of surface cleanliness – Part 1: Rust grades and preparation grades of uncoated steel substrates and of steel substrates after overall removal of previous coatings
ISO 8504-2: Preparation of steel substrates before application of paints and related products – surface preparation methods – part 2: abrasive blast-cleaning
ISO 12944-4: Paints and varnishes – corrosion protection of steel structures by protective paint systems – Part 4: Types of surface and surface preparation
ISO 12944-5: Paints and varnishes – corrosion protection of steel structures by protective paint systems – Part 5: Protective paint systems
ISO 20421-1: Cryogenic vessels – large transportable vacuum insulated vessels – Part 1: Design, fabrication, inspection and testing
ISO 21010: Cryogenic vessels – Gas/material compatibility
ISO 21013-1: Cryogenic vessels – Gas/material compatibility: Cryogenic vessels – pressure-relief accessories for cryogenic service – Part 1: reclosable pressure relief valves
ISO 23028: Cryogenic vessels – cleanliness for service (specify for LOX service)

## 1b)ISOtainer (product specific)
Product availability & associated specifications
Size, Tonnes LOX capacity | 10 | 20
Equivalent size available, litres
Boil-off rate (indicate conditions)
Discharge capacity (Nm3/hr)
Thermal conductivity
Insulation/fill material
Structural testing: wind & snow loading, earthquake rating

## 2) medical LOX manufacture
Component | Specification | Meets requirement? (Yes/No/NA) | Type of supporting document (select from drop down) | File name as submitted | Page No. | Comments
Volume | 20 tonnes (single delivery)
Regulatory | Proof of stringent regulatory authority (SRA) approval for medical LOX
OR
Proof of product meeting the requirements of an SRA pharmacopeia (e.g., USP, Ph Eu, BP, or IP) as shown by their specification and a sample Certificate of Analysis (CoA).
Good Manufacturing Practices(GMP) | Address of the manufacturing facility
Certificate of Analysis (CoA) template to indicate alignment with USP, Ph Eu, Int’l Pharmacopoeia, or British Pharmacopoeia
Copy of the manufacturer’s license to distribute medical oxygen or license as a wholesaler of medical oxygen
Proof of GMP OR the following:
• Certification of ASU, and storage tanks for oxygen service for medical application.
• Cleaning protocols for tanks (storage and transport) (e.g., cleaned to ISO 23208 with ‘cleaned for O2’ indicated in scope).
• Audit history of a manufacturing facility from the past 3 years (including validity dates for audits).
• Facility safety directives.
• If a manufacturer is producing industrial and medical LOX, an explanation of when the product streams are separated and how contamination is prevented.
• Explanation of how a batch of medical LOX is defined.
• Validation of manufacturing and cleaning processes.
Quality Management Systems | ISO 9001: Quality Management Systems - Requirements
Documentation | Evidence of sale of medical LOX to government body
Evidence of Long-Term Agreement with global procurement bodies (if available)

## 3) microVIE
Component | Specification | Meets requirement? (Yes/No/NA) | Type of supporting document (select from drop down) | File name as submitted | Page No. | Comments
Overview | Stainless steel or carbon steel outer shell
Stainless steel inner vessel (no aluminum)
Operating pressure: 1.4 bar (20 psi)
Fill source pressure requirement: 2.4-3.4 bar (35-50 psig)
Operating temperature: 10 °C /+40 °C
Able to be moved from ward to ward within health facility for gaseous oxygen delivery to patients
Sizing | Maximum LOX storage capacity: 45-60L
Features | Operating:
• Fill coupling / type for LOX (spec type available)
□ Top-fill line
• Gas withdrawal
□ Gas withdrawal vaporizer
□ Gas withdrawal valve
□ Economizer (preferred option)
□ Gas delivery rate: at least 15 LPM
• Pressure regulation
□ Pressure regulator with set-point feature for maximum allowable operating pressure (MAOP)
□ Line regulator for connection to gas outlet
• Flow control valve
• Overfill protection system
Monitoring:
• Liquid level gauge (& type to be specified)
Safety:
•Thermal relief valve
• Safety valves: primary (set at MAOP) and secondary (set at MAWP) (no bursting discs)
Additional offtake:
• Liquid withdrawal valve
Documentation | User and service manuals
Training materials or services offered on start-up, use for ISOtainer to microVIE filling, and preventive/curative maintenance
Regulatory | TPED 2010/35/EU or US DOT
FDA or CE under MDR
QMS | ISO 13485: Medical devices — Quality management systems — Requirements for regulatory purposes
Accessories | Roller base suitable for uneven or rugged terrain
Condensate bottle
1 set of transfill accessories including liquid oxygen transfer line, top fill line adaptor, vent wrench, and pressure gauge with DISS adaptor.
General design guidance (or equivalent) | EN 1251: Cryogenic vessels. Transportable vacuum insulated vessels of not more than 1 000 litres volume
Product Performance Standards (or equivalent) | ISO 18777: Transportable liquid oxygen systems for medical use — Particular requirements
ISO 23028: Cryogenic vessels – cleanliness for service (specify for LOX service)