## Instructions for specs
Please read the instructions carefully to ensure that all required documentation is submitted.
Only complete bids with all tabs filled and supplemental documentation attached will be considered.
Tab: "1 - LOT 1 Product Specifications"
This table contains a summary of the key documentation that must be submitted for a bid to be considered complete.
Please fill columns C, D, E as you prepare documents for your proposal and fill out the subsequent tabs of this spreadsheet.
For standards, please note that compliance to recent equivalents may be accepted upon review.
Tab: "2 - LOT 2 Product Specifications"
This table contains a summary of the key documentation that must be submitted for a bid to be considered complete.
Please fill columns C, D, E as you prepare documents for your proposal and fill out the subsequent tabs of this spreadsheet.
For standards, please note that compliance to recent equivalents may be accepted upon review.
Query period: | [insert dates]
Query email address: | ethiopiaprocurement@clintonhealthaccess.org

## Product Specifications
Product > | 1.Electrical Safety Analyzer |  |  |  | 2. Patient Monitor Simulator |  |  |  | 3. Gas Flow Analyzer and Ventilator tester |  |  |  | 4. Oxygen Analyzer |  |  |  | 5. Neonatal Incubator Analyzer |  |  |  | 6. Phototherapy Radiometers |  |  |  | 7. Electricians Tool Kit |  |  |  | 8. Phase sequence meter |  |  |  | 9. Digital Multimeter |  |  |  | 10. AC clamp meter |  |  |  | 11. Universal IC tester |  |  |  | 12. Multifunctional Soldering Station
Manufacturer >
Model >
Spec category | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments | Specifications | Meet requirement? | Type of supporting documents | Comments
Description | Electrical Safety Analyzer brings fast and simple automated testing in the form of a portable analyzer for biomedical and clinical engineering professionals that perform electrical safety testing on medical equipment both in the field and in facilities. | N/A | N/A | N/A | Multifunction patient simulator used as ECG simulator, fetal simulator, arrhythmia simulator, respiration simulator, temperature simulator, IBP simulator, cardiac output simulator, cardiac catheterization simulator, NIBP simulator, SpO2 tester biomedical technicians and engineers to check the performance of the vital sign monitoring equipment in the health facility. | N/A | N/A | N/A | Gas flow analyzers and ventilator testers are used by biomedical professional to test gas flow and respiratory medical equipment, including Ventilators and Anesthesia machine to measure Volume, pressure and flow. | N/A | N/A | N/A | Oxygen analyzer which is used by biomedical technicians and engineers for conformation of purity/percent levels of oxygen concentration, flow in the oxygen device (oxygen concentrator, Cylinder and plant). | N/A | N/A | N/A | Incubator analyzer which is used by biomedical technicians and engineers for conformation of incubator and radiant warmer sound levels and thermal characteristics, simultaneously measures airflow, relative humidity, sound, and four independent temperatures for maximum safe operation of infant incubator. | N/A | N/A | N/A | Phototherapy Radiometer which is used by biomedical technicians and engineers for the accurate detection and measurement of light radiation in the blue part of the spectrum. As an irradiance meter, it measures the output power of phototherapy units. | N/A | N/A | N/A | Electricians Tool Kit should have the following Package Contents | N/A | N/A | N/A | Phase sequence meter used for biomedical and electrical/technician /Engineers for teasing, measuring different electricity three phase sequence and motor rotation direction. | N/A | N/A | N/A | Digital multimeter used for biomedical technician /Engineers for teasing, measuring different electrical parameters. The meter should be equipped with power/ battery saving or auto battery off when the meter is not functioning. | N/A | N/A | N/A | AC clamp meter | N/A | N/A | N/A | Universal IC tester is for biomedical engineer or technicians used for testing different ranges of ICs in medical equipment maintenance workshops. | N/A | N/A | N/A | The station is a dual output unit supplied with soldering and desoldering tools. It should be microprocessor-controlled to allow independent temperature control of both tools. The station will automatically recognize the tool in use and calibrate itself for that tool. The station with a Vacuum for the desoldering tool is generated by an internal venturi driven from an external compressed air source. | N/A | N/A | N/A
Technical, Warranty, Regulatory, Quality | Leakage current: AC + DC, AC only, DC only |  |  |  | Types of tests: ECG (Including Fatal and arrhythmia), respiration, temperature, IBP, NIPB, cardiac output, SpO2 and rainbow multi wavelength wave form. |  |  |  | Measures of breath parameters: Inspiratory tidal volume, Expiratory tidal volume, Minute volume, breath rate, I:E,PIP, Inspiratory pause pressure, mean airway pressure, PEEP, Lung compliance, inspiration time , inspiration hold time, expiration time, expiration hold time, peak inspiration time and peak expiration time. |  |  |  | Measurement: oxygen purity, flow rate, outlet pressure |  |  |  | Display: not less than 2 x 16 super twist LCD |  |  |  | Measures: irradiance , spectrum range(wavelength) |  |  |  | 1. Utility storage box |  |  |  | Frequency range: 50 to 400Hz |  |  |  | Voltage DC Range: 4000 mV to 700V Accuracy: ±0.9% |  |  |  | AC current 600 A 0.1 A 2 % ± 5 digits |  |  |  | Tests a wide range of Digital IC's such as 74 Series, 40/45 Series of CMOS IC’s |  |  |  | Power Input: up to 165 W Temperature
| Mode of operation: Auto, Manual, ECG, Respiration simulation |  |  |  | Operating Temperature: 10°C to 32 °C |  |  |  | Gas corrections: ATP (ambient temp/pressure, actual humidity), ATPD (ambient temp/pressure, dry) Nitrogen (N2), ATPS (ambient temp/pressure, saturated), STP or STPD21 (21 °C temp/pressure 760 mmHg, dry), BTPS (body temp 37 °C/ambient pressure 760 mmHg, saturated) and BTPD (body temp 37 °C/ambient pressure 760 mmHg, dry) |  |  |  | Measured concentration range: 20.0% to 100% |  |  |  | Resolution: 1470 measurement points (210 records of seven parameters) |  |  |  | Spectral range: 429 - 490 nm (max. 97% response at 490 nm) |  |  |  | 2. Flat nose pliers 135 mm |  |  |  | Measurement parameter: Three phase rotation indication, Motor rotation indication |  |  |  | Voltage AC Range: 4000 mV to 600V Accuracy: ± (3%+3) |  |  |  | AC voltage 600 V 0.1 V 1 % ± 5 digits |  |  |  | can test Microprocessor 8085, 8086, Z80 |  |  |  | Control: Soldering/Desoldering iron continuously 50°C– 450°C
| Measurements: Earth resistance, leakage, ECG, respiration, DTU (Equipment under test) current, |  |  |  | Operating Maximum Humidity: 80 % relative humidity up to 31 °C, decreasing linearly to 50 % relative humidity at 45°C. |  |  |  | Types of tests: |  |  |  | Measured flow range: between 0-15LPM, Accuracy: ±2% |  |  |  | Power supply: 240 V input, with rechargeable battery |  |  |  | Measurement range: 0-199.99 µW/cm |  |  |  | 3. Bent nose pliers 130 mm |  |  |  | Voltage range: 40 to 7000V |  |  |  | Resistance Range: 100.0 Ω to 50.00 M Ω Accuracy: ± (1.25%+1) |  |  |  | DC voltage 600 V 0.1 V 1 % ± 5 digits |  |  |  | It tests Peripherals like 8255, 8279, 8253, 8259, 8251, 8155, 6264,62256,8288,8284 |  |  |  | Compressed Air: Inlet pressure 500 KPA– 600 KPA (58–87 psi) oil-free, dry compressed air Compressed Air
| Voltage Range (mains voltage): Single phase (220 Vac ± 10, 50Hz) and three phase (380Vac ± 10,50Hz) and Accuracy: ± 2 % |  |  |  | Altitude: Up to 2500m |  |  |  | Flow range: 0 to ±200 slpm (Standard litter per minute) , ±2% accuracy |  |  |  | Pressure range : 0.1-21 psi |  |  |  | Sound Level : 30 to 80 dBA ,Accuracy: ± 5 dB(A) |  |  |  | Resolution: 1 µW/cm |  |  |  | 4. Combination pliers 210 mm |  |  |  | Phase display: 120 to 400Vac |  |  |  | Capacitance Range: 1.000 µF to 10,000 µF Accuracy: ± (2%+2) Frequency: 20 Hz to 1 MHz |  |  |  | Resistance 6000 Ω 0.1 Ω 1 % ± 5 digits |  |  |  | tests Opamp,555, Transistor Arrays, Analog switches, Opto couplers and Others |  |  |  | Safety Class: 1 (control unit), 3 (soldering tool)
| Leakage tests: ground wire, chassis, lead to ground or to patient, lead to ground, lead to lead, lead to isolation, direct equipment, direct applied part, alternative equipment, alternative applied part and point to point. |  |  |  | Power: Rechargeable battery |  |  |  | Volume range: ±100 l , ±2% accuracy |  |  |  | Display: LCD screen |  |  |  | Display resolution: 0.1 dB(A) Should have at least four temperatures sensors Should be measure of mattress Convection (inside the incubator): - Three Sensors (T1, T3, T4) Convection (inside or outside the incubator): - One sensor (T2) |  |  |  | Probe: Lens matches the cosine receiving function of human skin |  |  |  | 5. Side cutting pliers 165 mm |  |  |  | Operation: Real time or continuous Operation |  |  |  | Battery: Alkaline, 650 hours continuous, Safety Complies with IEC 61010-1:2001or Equivalent, |  |  |  | Capacitance 10 μF to 1000 μF 1 μF 1.9 % ± 2 digits |  |  |  | tests Seven segment display of common cathode & common anode type |  |  |  | Foot switch connector: over adapter as accessories
| Insulation resistance test: range, source test voltage, insulation resistance test, |  |  |  | Battery Charger: with 240 V input, 15 V/2.0 A output. |  |  |  | High pressure range: -0.8 to 10 bar, ±2% accuracy |  |  |  | Resolution: 0.1% of output |  |  |  | Radiation: - One sensor for radiation head that can be used also for external temperature Temperature: reading value ± 0.5 ºC ± 1 LSB for the range from 25 ºC to 40 ºC |  |  |  | Power: 220Vac, 50Hz input and rechargeable battery with adopter |  |  |  | 6. Fine tip straight tweezers |  |  |  | temperature: 100C to 320C | [7 empty] | Able to have | NA |  |  | It has Auto search facility of IC’s 5 Test by: Truth table/sequence table comparison. |  |  |  | Certification: The product Should have CE/ISO or equivalent quality certification
| Earth Resistance Test current: > 200 mA ac Ranges: 0 Ω to 2 Ω Accuracy: ± 1 % |  |  |  | ECG performance test: |  |  |  | Differential low-pressure range: ±160 mbar, ±0.5 % accuracy |  |  |  | Response Time: ≤10S |  |  |  | Airflow: from 0.1 m/s to 2 m/s reading ± 0.1 m/s at temperature 25 ºC to 40 ºC and humidity 50 % RH ± 5 % RH, Display Resolution: 0.01 m/sec |  |  |  | Display: LCD screen |  |  |  | 7. Spring hook |  |  |  | Humidity: up to 85% | [7 empty] | · True-rms |  |  |  | ZIF: 40 pin DIP ZIF sockets. |  |  |  | The product Should have CE/ISO or equivalent quality certification
| Equipment Current: Range: 0 A to 20 A Accuracy: ± 5 % |  |  |  | Amplitude: 0.05mV to 0.5mV |  |  |  | Airway pressure range: ±160 mbar, ±0.5 % accuracy |  |  |  | Power Supply: 220Vac, 50Hz input and rechargeable battery with adopter |  |  |  | Operating and Storage temperature:10 °C to 40 °C, Humidity: 10 % to Up to 85 % non-condensing, Altitude: 2500 m |  |  |  | Certification: The product Should have FDA/IEC 60601-2-19/CE/ISO or equivalent quality certification |  |  |  | 8. Soldering aid tools (3 pcs.) |  |  |  | Power Supply : rechargeable battery | [7 empty] | · Continuity |  |  |  | Keys: 28 Touch-keys Key pad with numerical & functional keys. 5 Display: 9 Digit Seven Segment Display. |  |  |  | Warranty: One year warranty
| Leakage Current: Ranges: 0 µA to 5000 µA |  |  |  | Pulse wave: 30BPM, 60BPM |  |  |  | Barometric pressure range: 550 to 1240 mbar ±1 % accuracy |  |  |  | Sound Level: 30 to 80 dB |  |  |  | Weight with sensors 1.5 kg to 3kg (3 lb) |  |  |  | Warranty: Two-year warranty |  |  |  | 9. Disordering pump |  |  |  | Certification: The product Should have EN61010, EN6155-7,/CE/ISO or equivalent quality certification | [7 empty] | · Hold |  |  |  | Supply Input Voltage: 230V AC. |  |  |  | Documentation: Users and service manual in English Language should be included
| ECG performance: Ranges: 30 BPM to 250BPM and Accuracy: ± 1BPM |  |  |  | Wave type: Square wave(0.125Hz- 2Hz), Triangular wave(0.125Hz- 2Hz), Sine wave(0.05Hz to 150Hz) |  |  |  | Temperature measurement range: 0 to 50 °C |  |  |  | Sensor: Calibratable oxygen sensor |  |  |  | Communication: USB port and RS-232 cable recommended (sourced separately) |  |  |  | Documentation: Users and service manual in English Language should be included with calibration certificates and SOP. |  |  |  | 10. Drill set w/adaptor |  |  |  | Warranty: One year warranty | [7 empty] | · Backlight |  |  |  | Certification: The product Should have CE/ISO or equivalent quality certification
| Housing: Metal case or high impact plastic Should supplied with all necessary accessories. |  |  |  | Performance testing in ECG: QRS detection, T-wave rejection and R-wave detection. |  |  |  | Humidity measurement range: ±3 % RH (20 to 80 % RH) |  |  |  | Certification: The product Should have FDA/CE/ISO or equivalent quality certification |  |  |  | Certification: The product Should have FDA/CE/ISO/ IEC 61326-1, IEC 60601-2-19, IEC 60601-2-20, IEC 60601-2-21 or equivalent quality certification | [7 empty] | 11. Anti-static alignment tool kit (3 pcs.) |  |  |  | Documentation: Users and service manual in English Language should be included with certificate of calibration. | [7 empty] | Wire capacity 750 MCM |  |  |  | Warranty: One year warranty
| Operating temperature: 10 ºC to 32 ºC, Operating humidity 10 % to 85 % non-condensing |  |  |  | Lead selection: All ECG leads (LL,LA,RA,RL, V1,V2,V3,V4,V5,V6) |  |  |  | Oxygen: 0 to 100 %, ±3 % accuracy |  |  |  | Warranty: Two-year warranty |  |  |  | Accessories and spare parts: Should supplied with all accessories, Software, and spare parts. | [7 empty] | 12. Diagonal cutting nipper pliers | [11 empty] | The product Should have CE/ISO or equivalent quality certification |  |  |  | Documentation: Users and service manual in English Language should be included
| Certification: The product Should have IEC62353-60601-1/ANSI/AAMI ES-1/ISO or equivalent quality certification. |  |  |  | Pressure variabilities: Systolic and diastolic pressure are independently variable in 1 mmHg |  |  |  | Operating Temperature: 10°C to 32 °C |  |  |  | Documentation: Users and service manual in English Language should be included with calibration certificates and SOP. |  |  |  | Warranty: One year warranty | [7 empty] | 13. Long nose pliers | [11 empty] | Warranty: One year warranty
| Power requirement: 230 VAC/50 Hz, |  |  |  | Respiratory Rate: 0 to 150 BrPM |  |  |  | Operating Maximum Humidity: 80 % relative humidity up to 31 °C, decreasing linearly to 50 % relative humidity at 45°C. | [7 empty] | Documentation: Users and service manual in English Language should be included | [7 empty] | 14. Long nose pliers 150 mm | [11 empty] | Documentation: Users and service manual in English Language should be be included
| Humidity: Up to 85% non-condensing |  |  |  | Apnnea selection: 12 sec to 32 sec or continuous. |  |  |  | Altitude: Up to 3000m | [7 empty] | NB: Price quotation includes users and Maintenance training for three Biomedical Engineers | [7 empty] | 15. Adjustable wrench 6"
| Test recorded capacity: Up to 5000 |  |  |  | Measuring temperature: 30°C - 42°C with ± 1 % accuracy |  |  |  | Power supply: 240 V input, with rechargeable battery | [15 empty] | 16. Adjustable wrench 8"
| Warranty: Two-year warranty |  |  |  | Cardiac Output: 2.5 , 5, 10 LPM ± 7.5% accuracy |  |  |  | Interface: RS232 bi-directional interface/ USB | N/A | N/A |  |  | N/A | N/A |  |  | N/A | N/A |  |  | N/A | N/A |  | 17. Soldering iron25 Watt
| Documentation: Users and service manual in English Language should be included with calibration certificates and SOP. |  |  |  | NIBP test range: 80-300mmHg |  |  |  | Should supplied with all accessories. | [15 empty] | 18. Screwdriver 5 mm
| NB: Price quotation includes users and Maintenance training for three Biomedical Engineers |  |  |  | SpO2 test range: HR 30BPM to 300BPM in one BPM steps synchronized with ECG. |  |  |  | Housing: Metal case or high impact plastic | [15 empty] | 19. Screwdriver 6 mm
[5 empty] | Interface: RS232 bi-directional interface/ USB |  |  |  | Certification: The product Should have IEC /CE/CSA or equivalent quality certification. | [15 empty] | 20. Screwdriver 3.2x75 mm
[5 empty] | Should supplied with all accessories. |  |  |  | Warranty: Two-year warranty | [15 empty] | 21. Screwdriver #0x75 mm
[5 empty] | Housing: Metal case or high impact plastic |  |  |  | Documentation: Users and service manual in English Language should be included with calibration certificates and SOP. | N/A | N/A |  |  | N/A | N/A |  |  | N/A | N/A |  |  | N/A | N/A |  | 22. Screwdriver 5.0x75 mm
[5 empty] | Certification: The product Should have IEC61010-1 /CE/CSA or equivalent quality certification. Electromagnetic Compatibility (EMC): IEC or Equivalent | [19 empty] | 23. Screwdriver #1x75 mm
[5 empty] | Warranty: Two-year warranty | [19 empty] | 24. Screwdriver 6.0x100 mm
[5 empty] | Documentation: Users and service manual in English Language should be included with calibration certificates and SOP | [19 empty] | 25. Screwdriver #2x100 mm
[5 empty] | NB: Price quotation includes users and Maintenance training for three Biomedical Engineers | [19 empty] | 26. Screwdriver #2x157 mm
27. Disordering wick
28. Brush
29. Electronic screwdriver set (6 pcs.)
30. Socket and screwdriver set (40 pcs.)
31. Precision wire stripper
32. Crimping tool
33. Needle file set (5 pcs.)
34. IC extractor
35. Mini flashlight
36. Folding type hex key set (7 pcs.)
37. 3 Prong holder
38. Curved-claw hammer
39. PVC insulated tape
40. Solder (63%, SN)
41. Heat sink
42. Wrist strap
43. Measuring tape 3M
44. Electronic combination wrench (10 pcs.)
45. Inspection mirror
46. Slip joint pliers 154 mm
47. Aluminum tool case with Pallets
N.B All the equipment listed are brand and have ISO or equivalent certification

## RFI_MFLab (hidden)
| 关于产品 | [9 empty] | 关于贵司
产品类型 Product Type | 产品标准要求 Product Specifications | 产品型号（请列出贵公司符合左侧标准要求的产品，并邮件提供电子版产品用户手册） Product Model (s) | 产品排期（假设下周下订单，预计能够供的最早出货排期至何时（日期）？ Lead time if order placed next week (date) | 本次订单所需数量 Units Required | 公司当前可出售的数量 Units Available | 法规批准/认证Regulatory Approval/Certification |  | 产品性能标准Standards, for product performance |  | 产品单价（美金） Unit price (USD) | 制造商标准 Standards, for manufacturer |  | 预付款要求 Pre-payment requirement | 公司以前有向埃塞俄比亚出口的经验吗？Does company have previous export experience to Ethiopia? | 图片示例 Picture
[6 empty] | 标准 | 如果有，请填写 “X” | 标准 | 如果有，清填写 “X” |  | 标准 | 如果有，清填写 “X”
简单氧气气面罩 Simple Face Masks | 成人式： 医用级 | LB1110, Adult(LA), 5.5mm x 2m绿色管 |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X | 生产前付清 | 有
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
不重吸面罩 Oxygen Mask with reservoir bag | 成人式： 氧气浓度：支持80%至100％的氧气浓度 管长: 2m 使用标准接头；管路（氧气管线）不弯曲 材质：聚氯乙烯（PVC）； 透明；医用级 | LB1310, Adult(LA), 5.5mm x 2m绿色管 氧袋1000ml |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
文丘里面罩（氧浓度可调面罩） Venturi Mask | 成人式：能以8至10 L / min 的流速输送低至中等流量的氧气，特定氧浓度为30％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,31%(氧气流速6L/min±10%),白色，其他符合 |  | 300 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
| 成人式：能以10至12 L / min 的流速输送低至中等流量的氧气，特定氧浓度为40％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,40%(氧气流速8L/min±10%),粉色，其他符合 |  | 300 |  | CE | X | ISO 18190
[6 empty] | FDA | X | ISO18562-1 | X
[6 empty] | 出口许可证 | X | ISO/DIS 23368
| 成人式：能以12至15 L / min 的流速输送低至中等流量的氧气，特定氧浓度为60％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,60%(氧气流速8L/min±10%),红色，其他符合 |  | 300 |  | CE | X | ISO/DIS 17256 | X
[6 empty] | FDA | X | ISO 15223-1 | X
[6 empty] | 出口许可证 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 中等大小 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带, M |  | 130 |  | CE |  | ISO 18562-1:2017 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 17510:2015 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X |  |  |  | ISO14971 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 大码。 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带，L |  | 130
气管导管 Endotracheal Tube | 管： 尺码：7。 内径：7.0mm。 外径：9.3mm。 长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 袖口： 直径：24mm | LB5010C, 带囊，ID 7.0mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X |  | ISO14971 | X
| 管： 尺码：7.5 内径：7.5mm。长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 | LB5010C, 带囊，ID 7.5mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X